## Sheet1
Lab postcal comparision of FSI CT logger that was out at OCP1 for a year and retrieved Sep 2007.
The mesh net around the CT logger was removed by divers a few months into the deployment (Dan's memory…need a date),
or else it was deployed without the mesh from the start (Dave's records), but anyway, it is believed to have been reading high
without its mesh, since the net was present when it was calibrated.
This postcal comparison will allow back correction for the calibration offset (jump up in Sal) due to removal of the net mesh.
Raw Seawater (RSW) from the spigot and diluted RSW were used to compare readings of the postcal instrument
and a freshly calibrated CT logger which had not been deployed.
Comparisons were done in the middle of a carboy, on 10/2/2007
Good CT logger (never deployed) |  |  |  | Retrieved CT logger (for postcal): |  |  |  | Postcal offset correction
S/N 1606D-8Feb05 model CTS-C-1DH |  |  |  | S/N 1654D-09Dec05 model CTS-C-2DH
Cond | Temp | Sal |  | Cond | Temp | Sal |  | Cond
Raw Sea Water from spigot:
53.128 | 23.486 | 0 | 1st file, last capture val
|  |  |  | 51.06 | 23.353 | 34.7725 | 1st file, last val
53.16 | 23.554 | 0 | 2nd file, val chosen to match T | 54.153 | 23.553 | 36.9834 | 2nd file, last val | -0.993
53.153 | 23.584 | 0 | 2nd file, last val before removed
Conductivity offsets were different
at different Salinities,
RSW diluted with DI 50% | [8 empty] | so it's not a fixed offset but probably
32.795 | 23.64 | 0 | 1st file, last capture val | [5 empty] | proportional to salinitiy.
32.781 | 23.494 | 0 | 2nd file, val chosen to match T | 33.409 | 23.493 | 21.6191 | Only measured once | -0.628
32.781 | 23.59 | 0 | 2nd file, last capture val
Choose yellow highlighted vals for comparison: temps are similar
Notice I only did 2 salinities so I’m just making a line through 2 pts.
Assume zero intercept, ie assume that both CT loggers would have
given zero conductivity at zero salinity, and use the forced intercept
equation for correction:
Measured bad cond = 1.0188 * Correct Cond
so….
Corrected Cond = Conductivity measured/1.0188

## emails
Hey Xiaoyan,
I plotted up the data and it looks like the old CT logger was reading
0.08 mS/m high. There was a change from 5.65 to 5.57 mS/m.
Anybody else want to weigh in on this conversation?
Dave
David Wells
Coastal Ocean Research and Monitoring Program
UNCW Center for Marine Science
5600 Marvin K. Moss Lane
Wilmington, NC 28409
wellsd@uncw.edu
Lab (910) 962-2327
Cell (910) 231-2919
-----Original Message-----
From: Qi, Xiaoyan
Sent: Tuesday, October 02, 2007 2:01 PM
To: Wells, Dave; Tenney, Walter
Cc: 'Wendy L. Woods'; Kennedy, Clark
Subject: RE: back-correction note for Oak Isl SAL
Hi Dave,
Do you know whether conductivity was affected?
Thanks,
Xiaoyan
-----Original Message-----
From: Wells, Dave
Sent: Tuesday, October 02, 2007 9:40 AM
To: Qi, Xiaoyan; Tenney, Walter
Cc: 'Wendy L. Woods'; Kennedy, Clark
Subject: RE: back-correction note for Oak Isl SAL
Hey Xiaoyan,
The CT logger deployed from 8/06 to 9/07 was originally calibrated at
the factory with a small mesh screen on it. When the CT logger was
deployed in the field with this screen on, we found that it actually
"promoted" fouling of the instrument. For this reason, the decision was
made to remove the screen. So, when we removed the screen it caused a
0.5 PSU increase in salinity readings for that time period. The new CT
logger that is currently out there was calibrated without this screen,
so the readings are correct and fouling should not be an issue.
Dave
David Wells
Coastal Ocean Research and Monitoring Program
UNCW Center for Marine Science
5600 Marvin K. Moss Lane
Wilmington, NC 28409
wellsd@uncw.edu
Lab (910) 962-2327
Cell (910) 231-2919
-----Original Message-----
From: Qi, Xiaoyan
Sent: Tuesday, October 02, 2007 1:32 PM
To: Wells, Dave; Tenney, Walter
Cc: 'Wendy L. Woods'; Kennedy, Clark
Subject: RE: back-correction note for Oak Isl SAL
I get condfused again. Had the salinity data reported high by 0.5PSU
since 08/22/2006 or the salinity data slowly went lower during the
deploied time period?
Look at the salinity data below:
Year 20007:
http://152.20.240.9/universal2.php?station=OCP1&dtype=salinity_7.2&start
d=2007-09-24&endd=2007-09-27&timezone=GMT&format=graph&submit=Submit
Year 2006:
http://152.20.240.9/universal2.php?station=OCP1&dtype=salinity_7.2&start
d=2006-09-24&endd=2006-09-30&timezone=GMT&format=graph&submit=Submit
For the similar time period, 2006 reported lower value than 2007.
Comments, please!
Xiaoyan
-----Original Message-----
From: Wells, Dave
Sent: Tuesday, October 02, 2007 8:51 AM
To: Qi, Xiaoyan; Tenney, Walter
Cc: 'Wendy L. Woods'; Kennedy, Clark
Subject: RE: back-correction note for Oak Isl SAL
Hello all,
Based on dive logs, the CT logger was serviced on 7/27/06. This was
just after it was struck by lightning. It was not redeployed again
until 8/22/06. So, 8/22/06 should be the last time the CT logger was
serviced and the beginning of the time period that the CT logger
reported high by 0.5 PSU. I'm pretty sure that the salinity jump from
18 to 30 happened when we removed the screen from the CT logger at ILM1,
not OCP1. I plotted up the data from 4/06 to 11/06 at OCP1 and it never
got below ~28.
Thanks,
Dave
David Wells
Coastal Ocean Research and Monitoring Program
UNCW Center for Marine Science
5600 Marvin K. Moss Lane
Wilmington, NC 28409
wellsd@uncw.edu
Lab (910) 962-2327
Cell (910) 231-2919
-----Original Message-----
From: Qi, Xiaoyan
Sent: Tuesday, October 02, 2007 12:26 PM
To: 'Wendy L. Woods'; Tenney, Walter; Wells, Dave
Cc: Kennedy, Clark
Subject: RE: back-correction note for Oak Isl SAL
Hi All,
If any of you remember when the cleaning took place, please let me know.
Thanks,
Xiaoyan
-----Original Message-----
From: Wendy L. Woods [mailto:woodswendy@woodswendy.com]
Sent: Thursday, September 27, 2007 2:08 PM
To: Qi, Xiaoyan; Tenney, Walter; Wells, Dave
Cc: Kennedy, Clark
Subject: back-correction note for Oak Isl SAL
Dan tells me that there was a 0.5 jump-down in Sal when the CT sensor at
Oak Isl. was replaced yesterday. He says SAL was probably reading high
by a fixed offset for the last 1.5 yrs ever since the mesh netting was
removed during a dive to clean it. Since the mesh netting was present
when it was calibrated, removing it would have caused a fixed offset in
the calibration. At the time that the cleaning took place, about 1.5 yrs
ago, Dan remembers the Sal vals jumping from like 18 to the 30s, because
lots of fouling was removed at the same time. So it should be easy
enough to find out exactly what day that cleaning took place by the data
values, and perhaps Dave even has a record of which day they dove on
it? It would probably be worth to put a note somewhere indicating that
we believe the last 1.5 yrs (specific dates) of Sal data at Oak Island
are reading high by approx 0.5 ppt due to the fact that the mesh was
removed, so that potential users know to correct the data. Do we have a
place for such a note?
WW